## Slide 1
Electronic Pipette
Prototype v2 · Linear actuator
Precision Improvement via Linear Stepper Control
Comparative Analysis: Control · v1 · v2
Kahaan Padiya · June 2026

## Slide 2
What is v2 — Linear Step Parametrisation
🔩 Linear Actuator step Control
Replaces the Pulley system with force directors, with a industrial grade linear actuator, providing no backlash and much higher precision, Repeatability
📐 Automatic Calibration
Step parameters are configurable via serial interface. At startup, the pipette automatically homes using a contact switch at the bottom of piston stroke. Volume maps directly to step count, enabling reproducible and tunable aspiration depth.
🧪 Tighter Test Protocol
The readings are taken at 8uL, down from the 50uL, showing that the improvement in variation stays constant at lower volumes.
🎯 Target Improvement
v1 showed CV of 2.38% – mainly from string-slack at mechanical stops. v3 aims to bring CV below 0.8% by removing the string and using a direct screw drive.

## Slide 3
Raw Readings @ 520 nm — All Three Conditions
Direct absorbance values used for statistical comparison
Manual Pipette (Control)
Reading 1
0.47792
Reading 2
0.47687
Reading 3
0.47757
Reading 4
0.47685
Reading 5
0.47641
Reading 6
0.47592
n = 6
Electronic v1
Reading 1
0.54012
Reading 2
0.56028
Reading 3
0.56806
Reading 4
0.54440
n = 4
Electronic v3 (Linear Step)
Reading 1
0.57070
Reading 2
0.57237
Reading 3
0.57199
Reading 4
0.57524
Reading 5
0.56357
n = 5

## Slide 4
Precision Metrics — Control vs v1 vs v3
Metric
Manual (Control)
Electronic v1
Electronic v3 (Linear)
Average
0.47692
0.55322
0.57077
σ (Sigma)
0.000733
0.01316
0.00436
3σ
0.00220
0.0395
0.01307
CV %
0.154%
2.379%
0.763%
3×CV
0.461%
7.136%
2.289%
Readings taken at 520 nm · Pathlength 10 mm · Green = best, Red = worst CV

## Slide 5
CV % Comparison — Improvement Across Versions
Chart: CV (%)
| Manual Control | Electronic v1 | Electronic v3 (Linear)
CV % | 0.15379 | 2.3786 | 0.76305
Control CV
0.154%
v1 CV
2.379%
v3 CV
0.763%
v3 is 68% better CV than v1

## Slide 6
Individual Reading Spread — All Conditions
Chart: Absorbance @ 520nm
| R1 | R2 | R3 | R4 | R5 | R6
Manual Control | 0.47792 | 0.47687 | 0.47757 | 0.47685 | 0.47641 | 0.47592
Electronic v1 | 0.54012 | 0.56028 | 0.56806 | 0.5444
Electronic v3 (Linear) | 0.5707 | 0.57237 | 0.57199 | 0.57524 | 0.56357

## Slide 7
Analysis — What v3 Tells Us
✅ CV cut from 2.38% → 0.76%
Linear step control dramatically reduced inter-reading variability. The 68% improvement over v1 confirms that string-slack at mechanical stops was the dominant source of error.
⚠️ Still 5× above manual baseline
Control CV = 0.154%. v3 at 0.76% is a major step forward but further work remains. Likely sources: Flex in 3d print, loss of motor power due to sideways torque
📊 Volume control implemented
Through weighing scale calibration, we can now correlate volume with step count, and control the volume via the app interface
🔧 Next steps to close the gap
Adding a pressure switch for self calibration

## Slide 8
Conclusion
Electronic Pipette v3 — Linear Step Parametrisation
✅ CV reduced from 2.38% (v1) → 0.76% (v3) — a 68% improvement
✅ Linear step control eliminates stop-dependency and string-slack errors
✅ 3σ band tightened from 0.0395 (v1) to 0.0131 (v3) absorbance units
✅ Volume control enabled using step count and calibration with weighing scale